# BASIC  OPERATIONS

(Newest manual at:	 http://chemiris.chem.binghamton.edu/staff/schulte/AM-Tecmag-Manual.doc

Processing software:	 http://chemiris.chem.binghamton.edu/staff/schulte/ntnmr.exe

Alternative:		 http://chemiris.chem.binghamton.edu/staff/schulte/MestreC.exe )

## A.I	Locking and Shimming the Sample (on the console panel)

- Log on and immediately launch NTNMR.
Click the OK button and switch back to the black Lock Display window.

- Remove the cap from the magnet, press the orange button and then the “Lift” button.

- Remove the standard sample tube and replace it with your own cleaned sample tube. 
Press the “Lift Off” button to lower your sample into the magnet. Wait until the sample 
spins at 10-20 Hz and keep the temperature at 300 K during the experiment. Cap the magnet.

- Set the “LOCK POWER” and the “FIELD” to the proper values for your solvent:

| Solvent | LOCK POWER | FIELD |  | Solvent | LOCK POWER | FIELD |
| --- | --- | --- | --- | --- | --- | --- |
| acetone-d6 | 25* | 550 |  | methanol-d4 | 30* | 640 |
| acetonitrile-d3 | 30* | 560 |  | pyridine-d5 | 30* | 950 |
| benzene-d6 | 25* | 900 |  | TFA-d | 45* | 1200 |
| CDCl3 | 40* | 900 |  | THF-d8 | 40* | 500 |
| CD2Cl2 | 35* | 750 |  | toluene-d8 | 30* | 900 |
| DMSO-d6 | 25* | 550 |  | water-d2 | 30* | 700 |

* For concentrated samples only (more than 10 % or more than 100 mg of sample):  
 		Raise the LOCK POWER setting by 5 to 10 units.

- Refine the “FIELD” until you can cleary see the lock ringing signal.  
If you use two solvents, try both FIELD settings and use the stronger signal.

- Set the “LOCK PHASE” to 250, then press the “AUTO LOCK” button and wait until its diode and the  “LOCK GAIN” diode have stopped flashing.
This may take anywhere from a few seconds to several minutes.
Now the lock signal should look like this:
Choose between automatic or manual shimming (see next page).

## A.Ia	Manual Shimming of the Sample (using the black lock window and the wheel)

- Is the lock signal off screen? Decrease “LOCK GAIN” until you can see it. (Minimum is 22).
Is the lock signal still too high? Decrease the “LOCK POWER” until you can see it.
Is the lock signal too noisy? Increase the “LOCK POWER” and decrease the “LOCK GAIN”.

- Adjust the “LOCK PHASE” to raise the position of the lock signal on the screen.  
The lock signal may go off the screen. That’s fine, just decrease the LOCK POWER or LOCK GAIN to let it reappear and then continue improving the LOCK PHASE.

- Adjust the “Z” and “Z2” shims individually to achieve the highest position for the lock signal. 
Again, use LOCK POWER or LOCK GAIN to bring back the signal, if it disappears off screen.

- Continue refining “Z” and “Z2” several times until there is no further improvement.

- Press the “STDBY” button to deactivate the wheel. You can run your experiments now.
(You may be able to improve the shims by trying automated shimming.)

## A.Ib	Automated Shimming

- Switch to the NTNMR program and open the file “1 - AdjustShims.tnt”.

- Click the Console Toolbar button (button #3) and make sure the shims section is displayed.
In the Coarse Shims section check only the Z1 and Z2 boxes. All others must remain unchecked.
You may change the Z1 and Z2 values, if you remember good settings from previous samples. 
In the Shim Parameters section set:
 		Delay: 2s		Step: 10		Target: 0.01

- Click on the green ZG button. This will run one single scan (3 seconds).
If the FID is more than 1 cm high, select FID Shim.
If the FID is less than 1 cm high, select Lock Shim.

- Click “Read” and then “Start” at the bottom of the window to start automatic shimming. 
This procedure will maximize the signal strength by improving the Z1 and Z2 shims.
The computer may take several minutes to complete shimming your sample. 
The “Start” button will be off (grey) while shimming. It will come on (black) when finished. 
Don’t do anything else while shimming is in progress! Don’t open any spectra or programs!
After shimming is finished turn off the console toolbar (button #3) and run your experiments.

For overnight 13C experiments only:  Press “AUTO SHIM”, “Z”, “Z2”, “AUTO SHIM”.
Do not do this for 1H NMR. It will broaden your peaks and may wipe out small H,H-couplings.

## A.II	Routine 1H  and 13C Experiments

- Click File | Open or press Ctrl-O
 	and select the desired parameter file from the directory C:\NTNMR\data:
 		Proton NMR parameters: 	 “1H – solvent.tnt”  or  “1 – Proton”
 		Carbon NMR parameters: 	“13C – solvent.tnt”

- Click on “File | Save as “  (not “Save”!!!!), then move to your research group’s subdirectory.

Type a NEW filename including your lab notebook page number.
This file will later be used to store your spectrum. NEVER OVERWRITE ANY OLD FILES!
Don’t save your files directly to a flash drive. Copy them later.

- Click the first button on the toolbar and type the sample name in the comment box.

- Click the second button on the toolbar or select “View | Dashboard “ to change parameter settings.
For dilute samples increase the setting for “Scans 1D” (Acquisition). Use multiples of 8 !!! 
For very concentrated samples (>100 mg) set the “Receiver Gain” to “1” (Hardware tab).

- Click on the green ZG button or press Ctrl-Z to start the data acquisition.
The	     button will display a real-time view of the spectrum. 
 		     This will allow you to see, whether the sample is very dilute (increase Scans1D) 
 		     or too concentrated (If the baseline is distorted, reduce the Receiver Gain to “1”). 
The status bar at the bottom will indicate the progress of the experiment. Let the scans finish, or:
When the signal-to-noise is sufficient, you can stop the experiment with the yellow stop button. 
If you run an overnight experiment, leave a note with instructions for stopping and saving the data.

- Press Ctrl-S. Only this step will save your data.

- Create a new file with ”File | Save as” and add the word “processed” to the current filename.
Use only this file for processing. You can simply “Save” after each of the next processing steps.

- Press Ctrl-B to correct the baseline.

- 13C only: press Ctrl-E for an exponential multiplication. (also for 1H of extremely dilute samples)

- 1H only: Use “Commands | Zero Fill Commands | Zero Fill” to improve the resolution.

- All data: Press Ctrl-F to Fourier-transform the data and Ctrl-H for an automatic phase correction.

- Perform a manual phase correction (if needed):
Select “Option | Phase Adjustment ”, click on the strongest peak and click the Set Pivot button. 
Use the “Phase 0” slider to adjust the phase of the signal at the pivot point. 
Use “Phase 1” to adjust the other signals in the spectrum. 
If you reach the end of the range for either slider, click the Apply button below the sliders and continue adjusting the phase. Finally, click the other Apply button at the bottom when finished.

- Only after the phase correction is perfect  try a manual baseline correction:
(only necessary if the baseline isn’t flat) 
Display the complete spectrum and use the arrow keys to increase the baseline to ½ inch. 
Select “Option | Baseline Fix ” and click the ”Auto” and “Calculate” buttons.
Match the curvature of the baseline as closely as possible by moving the        boxes. 
Add baseline points with the right mouse button if necessary. Click “Apply”.

- Change the appearance of the ppm scale (This is only a one-time-adjustment for a PC): 
Double-click on the ppm scale below the spectrum to bring up the Axis Setup window. 
My recommended settings in the “Tick Marks” box are: 1, 5 and 1.
Uncheck “Show Grid” and check “Adjust Ticks on Zoom”. Set “Units” to “ppm”.

- Zoom functions:
  You can rescale the signal intensities with the “Fit To Window” button.
Double-click on the mini-spectrum on the top to return to the full width.
  The Up and Down arrow keys on the keyboard adjust the vertical size of the spectrum.
You can zoom into the spectrum with a click-and-drag operation of the mouse.

- Calibration of the chemical shift reference (use the solvent signal or TMS, if present):
Expand the reference signal as much as possible and then left-click on its central peak.
Right-Click anywhere in the spectrum and select Processing | Set Reference.
Select ppm units and enter the correct reference value. Click OK when done.

- Peak picking
Select Option | Peak Pick. 
You can adjust the height of the black horizontal bar that will appear.
All peaks that exceed the threshold of the black area will be picked. 
Click on Apply when done. The program will only number the peaks at this point. 
Zoom into the spectrum and delete or add peaks. Click Exit when you are done.
Later, before you print, you can decide what the labels should be (i.e. ppm).

- Integration
Select Option | Integrals. 
First, display the complete spectrum and click on Remove to erase any pre-existing integrals.
Highlight the first signal you wish to integrate. Make sure that you select a wide enough region 
(3-4 times the signal width) to measure accurate integrals. Expand regions as needed.
Right-click inside the highlighted area and select Add Integral. Repeat for other peaks.
Click on the Zero All button on the left panel.
Double-click on one of the integrals. (You should know how many protons it corresponds to.)
You can change the Active Integral to select the integral you wish to normalize. 
Set the Assigned value to the desired value. Click Apply and then Close. 
In the left panel change the labels to Assigned Value and Exit when done.

- Select Option | Amplitude Adjustment to define the width of the print region (F1 and F2).
Print each 1H spectrum from 10 to -0.5 ppm, 13C from 210 to -10. Print expansions if necessary.
(Your group may have different requirements for the printing ranges.)

- Use the File | Print Layout dialogue to select and modify the settings for the printout.
Uncheck Parameters, Line Fit, Print Frame, and Logo. Select an appropriate Pen Width.
Check Integrals and under Settings check Show Value. Use the same pen width as above.
Check Peak Picks and under Settings set Precision to 2 or 3 and select Label by chemical shift. 
Grab the ppm scale and drag it to the bottom of the red box.
The print screen key on the keyboard will export the complete screen to the clipboard. (Reports!)
Click on the printer icon to print out the spectrum. Only exit after printing!

## B.	Special 1D experiments:

## B.I	13C DEPT-45-90-135	to determine CH multiplicities.

## This experiment will record 3 13C spectra:
#1 (DEPT-45)   will have CH3, CH2, and CH peaks positive
#2 (DEPT-90)   will have only CH peaks
#3 (DEPT-135) will have CH2 negative and CH3 and CH positive.
Often times the DEPT-135 or the PENDANT experiments are sufficient, 
 	  see experiments B.II and B.III on the next page.

- Press Ctrl – O and open the file “DEPT-45-90-135.tnt”.

- Open the dashboard and set “Scans1D” to any multiple of 32.
(The experiment cannot be interrupted. Use the exact number of scans!
  32 scans will take approx. 8 minutes, 160 scans = 40 minutes.)

- In the frequency section of the dashboard: The F1 Offset Freq. for your solvent is:
D2O: 6 kHz;	Acetone, Acetonitrile, DMSO, Methanol, THF:	7.3 kHz

Chloroform, Benzene, Methylene Chloride, Pyridine:	5.5 kHz

- Use “File | Save as” to define the filename in your data directory.

- Start the experiment with Ctrl – Z.

- Press Ctrl – S to save the data after the third experiment has finished.

- Processing:
Ctrl – B, Ctrl – E, Ctrl – F, Ctrl – H to process the data on the screen.
If a manual phase adjustment is needed, use Option – Phase Adjustment.
Calibrate the spectrum using a known signal from your C-13 spectrum.
Click File – Revert to Saved.
Hold down the shift key and click the NDFT button.
Check the boxes:
 	Use 1D Settings, 	Baseline Correction, 		Fourier Transform, 
 	Use System Phase, 	Apodization: Exponential, 	LB: 2
Uncheck every box under the 2D heading.
Click Do It.

- Viewing:	Your toolbar should show you a few additional fields:
You can select the 2D trace (= the individual DEPT spectrum) you wish to see.

- Printing:	Turn off the grid under the View menu.
Zoom into the spectrum to expand the region you want to print. Adjust the height of the signals.
Go to File | Print Layout and pull down the top edge of the spectrum to 1/3 its original height.
Use Ctrl – Tab to return to the spectrum and switch to the second trace.
Go to File | Add to Print Layout and align the two spectra. Resize, if necessary.
Use Ctrl – Tab to return to the spectrum and switch to the third trace.
Go to File | Add to Print Layout and align the three spectra. Resize, if necessary.
Remove the axes and everything else you don’t want for the two top spectra.

- Click the Print button.

## B.II	13C DEPT-135 (negative CH2 peaks, CH3 and CH positive, no quat. C)

- Press Ctrl – O and open the file “DEPT.tnt”.

- Open the dashboard and set “Scans1D” to any multiple of 32.
(The experiment cannot be interrupted. Use the exact number of scans you need!
  32 scans will take approx. 2 minutes – approx. 1000 scans per hour.)

- Use “File | Save as” to define the filename in your data directory.

- Start the experiment with Ctrl – Z.

- Press Ctrl – S to save the data after the experiment has finished.

- Processing just like for normal C-13 spectra.
(Phasing: Remember that CH2 signals are negative, CH and CH3 signals positive.)

## B.III	13C Pendant	(same as DEPT135, additionally quat. Carbons will be negative)

- Press Ctrl – O and open the file “Pendant.tnt”.

- Open the dashboard and set “Scans1D” to any multiple of 32.
(The experiment cannot be interrupted. Use the exact number of scans you need!
  32 scans will take approx. 2 minutes – approx. 1000 scans per hour.)

- Use “File | Save as” to define the filename in your data directory.

- Start the experiment with Ctrl – Z.

- Press Ctrl – S to save the data after the experiment has finished.

- Processing just like for normal C-13 spectra.
(Phasing: Remember that Cq and CH2 signals are negative, CH and CH3 signals positive.
  	    The solvent signal will also be negative, TMS - if present - will be positive.)

## B.IVa	13C Gated Decoupling	(allows measurement of C,H coupling constants)

- Press Ctrl – O and open the file “Gated Decoupling.tnt”.

- Open the dashboard and set “Scans1D” to any multiple of 8.
(The experiment needs at least as many scans as the regular C-13 experiment.)

- Use “File | Save as” to define the filename in your data directory.

- Start the experiment with Ctrl – Z.
(If you try to stop the experiment, it may continue to run for several minutes! Don’t panic!)

- Press Ctrl – S to save the data after the experiment has finished.

- Processing just like for normal C-13 spectra.

## B.IVb	13C Inverse Gated Decoupling	(allows integration of 13C signals)

- This is an overnight experiment and requires a very concentrated sample.

- Press Ctrl – O and open the file “InverseGatedDecoupling-C-13.tnt”.

- Open the dashboard and set “Scans1D” to any multiple of 8.
Set the “Last Delay” to “30s”.

- Use “File | Save as” to define the filename in your data directory.

- Start the experiment with Ctrl – Z.

- Press Ctrl – S to save the data after the experiment has finished.

- Processing just like for normal C-13 spectra.

Water signal suppression

## B.V	1H NOE Difference Measurements for one Proton

Finding the signals for irradiation:

- Shim your sample using Z and Z2, then turn off sample spinning! (VERY IMPORTANT!)

- Shim your sample using X and Y. Keep the spin off during the whole experiment.

- Read the 1H parameter file for your solvent.

- Open the Dashboard, click on the Frequency tab, and write down the “F1 Offset Freq.”
Set the “F2 Offset Freq.” to zero. 		Keep the dashboard open.

- Start ZG to get your initial proton spectrum.

- Process with: Ctrl-B, Ctrl-F, Ctrl-H. No need for perfect phasing.

- First, click the mouse pointer on a point in the spectrum where there are no signals (~15 ppm).
Select Scripts | Acquisition Scripts | Set Dec. Freq.
If the F2 Offset Freq. is still “0.000… kHz”, restart from step 3.
Write down the value of the new “F2 Offset Freq.” from the dashboard.

- Then, click the mouse pointer on the signal which you want to irradiate.
There should be no other signals within ½ ppm.
Select Scripts | Acquisition Scripts | Set Dec. Freq.

Write down the value of the new “F2 Offset Freq.” from the dashboard.

Setting up and starting the NOE experiment:

- Open the File: “NOE Difference for one proton.tnt”

- Enter the “F1 Offset Freq.” that you found in step 2.

- On the Dashboard, set the “F2 Offset Freq.” to zero.
Set “Scans1D” to the desired number of scans (should be several hundred).
(800 scans will take approx. two hours.)

- Click on the Pulse Sequence icon to open the pulse sequence editor.
In the line O2: right-click on the red icon to open the frequency table.
Enter both frequency values you wrote down into the table.

- Close the pulse sequence editor.

- Define a filename (File | Save As).

- Start the experiment with: ZG.
Don’t forget to save the data again after the experiment has finished.

Processing and evaluation:

- Ctrl-B, Ctrl-E, Ctrl-F, Ctrl-H
Correct the phase to make the irradiated signal negative and the NOE signals positive.
Integrate the NOE signals and the negative reference peak, set the reference integral to “-100”.
The integrals show directly the NOE enhancements in percent.

- Print using the Print Layout.
Print using the same ppm range as for the corresponding proton spectrum.

## B.VI	1H NOE Difference Measurements for more than one Proton
This experiment generates one large file which has to be split into 4 or 8 pieces.  
 Therefore you have to define 4 or 8 frequencies (1 reference + 3 or 7 signals).

- Shim your sample using Z and Z2, then turn off sample spinning! (VERY IMPORTANT!)

- Shim your sample using X and Y. Keep the spin off during the whole experiment.

- Read the 1H parameter file for your solvent.

- Open the Dashboard, click on the Frequency tab, and write down the “F1 Offset Freq.”

- Open the file: “NOE Difference (CDCl3).tnt”
Enter the “F1 Offset Freq.” that you found in step 2.

- Start the Experiment to get your initial proton spectrum. Stop after 8 scans.

- Process with: Ctrl-B, Ctrl-F, Ctrl-H

- First, click the mouse cursor on a point in the spectrum where there are no signals.
Select Scripts | Acquisition Scripts | Set Dec. Freq.
Write down the value of the new “F2 Offset Freq.” from the dashboard.

- Then, click the mouse cursor on a signal which you want to irradiate.
Select Scripts | Acquisition Scripts | Set Dec. Freq.
Write down the value of the new “F2 Offset Freq.” from the dashboard.

- Repeat step 7 for all peaks you want to irradiate.
You need either 3 or 7 additional frequencies for this experiment to work.

- On the Dashboard, set the “F2 Offset Freq.” to zero.
Set “Points2D” to the number of frequencies that you have selected.
Set “Scans1D” to the desired number of scans (should be several hundred).
(800 scans will take approx. one hour per selected frequency.)

- Click on the Pulse Sequence icon to open the pulse sequence editor.
In the line O2_2D: right-click on the red icon to open the frequency table.
Enter all the frequency values you wrote down into the table (4 or 8).
(There have to be exactly as many entries in the table as you specified in “Points2D”.)

- Close the pulse sequence editor, define a filename (File | Save As) and start ZG.
Don’t forget to save the data again after the experiment has finished.

- Processing: Open the data file.

- Select Scripts | Processing Scripts | 2D to 1D and convert the data.

- Extract the single FIDs by using Commands | Data Manipulation | Read First Half
and/or Commands | Data Manipulation | Read Second Half combinations.
(i.e. for 4 spectra: #1 is first,first, #2 is first,second, #3 is second,first, #4 is second,second)
Save each segment separately using unique filenames.

- Reference spectrum: Read, process and print the reference spectrum (#1) first.
Generation of difference spectra: First, close all open spectrum windows!!!!!
Select Scripts | Processing Scripts | AddSubtract and select the files that you want to subtract. Make sure the Subtract button is activated. File2 always has to be the reference FID.
After finishing the subtraction use Ctrl-Tab to find the difference FID (smallest amplitude)
and process it using Ctrl-B, Ctrl-E, Ctrl-F, Ctrl-J, Ctrl-B. Integrate the NOE signals and the negative reference peak, set the reference integral to “-100”.
Print using the same ppm range as for the reference spectrum.
The integrals show directly the NOE enhancements in percent.

## B.VII	T1 and T2 Relaxation Measurements

- Read the parameter file for your solvent.

- For proton experiments, write down the “F1 Offset Freq.” from the dashboard.

- Open one of these files: T1-1H, T1-13C, T2-1H or T2-13C  for the experiment you want.

- For proton experiments, enter the “F1 Offset Freq.” from step 2.

- Set “Points2D” to “10”.
Set “Scans1D” to the desired number of scans (should be very small).
(1 scan will take approximately 2 (1H) or 4 (13C) minutes per delay.)

- Click on the Pulse Sequence icon to open the pulse sequence editor.
(T1) In the line Delay_2D: right-click on the red icon to open the list.
(T2) In the line Loop_2D: right-click on the red icon to open the list.
Enter the following values into the table:

1H-T1: 	0.01s  0.1s  0.5s  1s  1.5s  2s  4s  10s  25s  60s
13C-T1: 	0.01s  1s  3s  5s  8s  12s  15s  30s  60s  120s	
13C and 1H-T2: 2  20  50 100  200  300  400  500  750  1000		
(There have to be exactly as many entries in the table as you specified in “Points2D”.)

- Close the pulse sequence editor, define a filename (File | Save As) and start ZG.
Don’t forget to save the data again after the experiment has finished.

- Processing: Open the data file and process it as usual.
For the T1 experiments correct the phase until all peaks are negative. 
For the T2 experiments correct the phase until all peaks are positive. 
Then click on File | Revert to Saved.

- Hold down the Shift key and click on the NDFT button to open the processing dialogue.

- Under the 1D tab check: Use 1D settings, Baseline correction, Fourier Transform, and 
Use System Phasing. Apodization: Exponential with LB: 0.5 for 1H and 4 for 13C.
Click on Do It when done.

- Click on the 2D display button (sixth button from left).

- Right-Click anywhere in the spectrum, select Data Type | Real

- Set Mode to Stack Plot and adjust the levels to see all spectra.
Grab and move the edges and corners of the frame to define the appearance of the spectra.
Change the amplitude of the spectra as desired.

- Select File | Add To Print Layout and define the printout settings.

## B.VIII	Pulse Width Determination:

- Read the parameter file for your solvent. (Use a very concentrated sample.)

- Record one scan and process with Ctrl-B, (Ctrl-E for C-13 only), Ctrl-F, Ctrl-H.
Save the spectrum with File | Save as.

- Click the mouse cursor on a signal in the center of the spectrum.
Select Scripts | Acquisition Scripts | Set Obs. Freq.

- Open the Dashboard 
Under Acquisition set Points2D to “25”. Set Scans1D to “1”.Set Dummy Scans to “0”.
Under Sequence set h90 or p90 to “0u”. Set Last Delay to “60s”.

- Click on the Pulse Sequence icon to open the pulse sequence editor.
In the Delay line: click on h90 or p90 and select 2D Table.
A new event line (Delay_2D) will appear at the bottom.
Right-click on the de0:2 item in that line. A delay table box will open.
Select Auto and enter “2u” as increment value. Click OK.

- Close the pulse sequence editor, define a filename (File | Save As) and start ZG.
Don’t forget to save the data again after the experiment has finished.

- Processing: Open the data file and process it as usual. Manual phase correction!!!
Then click on File | Revert to Saved.

- Hold down the Shift key and click on the NDFT button to open the processing dialogue.

- Under the 1D tab check: Use 1D settings, Baseline correction, Fourier Transform, and 
Use System Phasing. Apodization: Exponential with LB: 0.5 for 1H and 4 for 13C.
Click on Do It when done.

- Click on the 2D display button (sixth button from left).

- Right-Click anywhere in the spectrum, select Data Type | Real

- Set Mode to Stack Plot and adjust the levels to see all spectra.
Grab and move the edges and corners of the frame to define the appearance of the spectra.
Change the amplitude of the spectra as desired.
Double-Click on the ppm scale and turn it off.

- Select File | Add To Print Layout and define the printout settings.

## B. IX	Decoupling Experiments (for CHEM 485f/585f only)

- Press Ctrl-O or click File | Open
 	and select the desired 13C parameter file from the directory C:\NTNMR\data:
 	Carbon NMR parameters: “13C – solvent.tnt”

- Open the Pulse Sequence Window and the Parameter Dashboard.
Under “Acquisition” set “SCANS_1D” to “8”.  Under “Sequence” set “Last Delay” to “5”. 
The following parameters have been optimized for 80% ethylbenzene in benzene-d6:

  - High Power CPD Decoupling: 	Under “Hardware” set “Dec. Attn.” to “116”.
					Insert “DEC_RF” pulses in every column.
					Insert and continue “Async Table” phases in the “DEC_PH” 
 					line for all columns, beginning with the second column.

  - High Power BB Decoupling: 	Under “Hardware” set “Dec. Attn.” to “116”.
 					Remove all phases from the “DEC_PH” line.
 					Insert “BB_MOD” pulses in every column.

  - Low Power BB Decoupling:		Under “Hardware” set “Dec. Attn.” to “140”.
 					Remove all phases from the “DEC_PH” line.
 					Insert “BB_MOD” pulses in every column.

  - Off-Resonance CW Decoupling:	Under “Hardware” set “Dec. Attn.” to “116”.
					Remove all pulses from the “BB_MOD” line.
 					Remove all phases from the “DEC_PH” line.
 					Insert “DEC_RF” pulses in every column.
 Run off-resonance experiments for each of the following five “F2 Offset Freq.” values:
 	 4.408 kHz (CH3 on resonance), 4.889 kHz (CH2 on resonance), 
 	 6.566 kHz (aromatic H’s on resonance), 10 kHz, 20 kHz.

- Press Ctrl-Z or click on the ZG button to start the data acquisition. 
(Wait for the experiment to finish. The counter on the bottom right will indicate the progress.)

- Click on “File | Save as “  (not “Save”!!!!), then move to your group’s subdirectory, 
 	and select a NEW filename which you want to use to store your spectrum.

NEVER  OVERWRITE  ANY  OLD  FILES!

- Repeat from Step 2 to perform the rest of the experiments in a similar manner.

- Processing can be done offline on any computer with processing software. 
Process as any C-13 spectrum.
Print spectra 2a and 2b with a ppm scale and ppm peak picking.
Print all other spectra with a Hz scale and Hz peak picking.

## B. X	Water Signal Suppression

- Press Ctrl-O or click File | Open
 	and select “solvent suppression.tnt”

- Click on “File | Save as “  (not “Save”!!!!), then move to your group’s subdirectory, 
 	and select a NEW filename which you want to use to store your spectrum.

NEVER  OVERWRITE  ANY  OLD  FILES!

- Open the Parameter Dashboard.
Under “Acquisition” set “SCANS_1D” to any multiple of 8.

- Start ZG to get your initial proton spectrum.

- Immediately stop the experiment with the yellow stop button and wait for the scans to finish.

- Process with: Ctrl-B, Ctrl-F, Ctrl-H. No need for perfect phasing.

- If the water signal is still very strong, click on it.

- Select Scripts | Acquisition Scripts | Set Dec. Freq.

- Click on the green ZG button or press Ctrl-Z to start the data acquisition again.

The	     button will display a real-time view of the spectrum. 

This will allow you to see, whether the solvent signal (HDO) has been suppressed. 
If it is still too strong, the “F2 Offset Frequency” needs to be adjusted again from step 5.

- Press Ctrl-S to save your data.

- Create a new file with ”File | Save as” and add the word “processed” to the current filename.

- This FID can be processed like any other proton spectrum.

# C.	Special 2D experiments: (Additional training is needed for your first sample!)

## C.I	COSY / NOESY (Durations with default settings: approx. 1 hour / 10 hours)

- Press Ctrl – O and open a Proton NMR file.

- Record a proton spectrum, Fourier-transform, calibrate and save it (Ctrl – S).

- Display the complete spectrum on the screen.
Are there any signals above 9 ppm or below -1 ppm? (This is important for step 5!)

- Press Ctrl – O and open the file “COSY.tnt” or “NOESY.tnt”.
Use “File | Save as” to define the filename in your data directory.

- Select:	View | Dashboard and select the Acquisition tab.
Make Points1D and Acq.Points both 1024.
Make Points2D = 512 and Scans1D = 8 (for NOESY use 32).
If there were any signals above 9 ppm or below -1 ppm in your proton spectrum 
you will have to increase SW+/- and SW2D by 0.7 kHz for every additional ppm.

- Switch to the Frequency tab.

- Make F1 Base Freq. and F2 Base Freq. both 360.13.
Each solvent needs its own values in the F1 Offset Freq. and the F2 Offset Freq. fields:
(incorrect settings will lead to undesireable folding effects!!)
Acetone:	7.3 kHz		Methylene Chloride:	6.1 kHz
Acetonitrile:	7.6 kHz		Pyridine:		5.4 kHz
Benzene:	5.5 kHz		Trifluoro acetic acid:	4.7 kHz
Chloroform:	5.5 kHz		Tetrahydrofuran:	6.7 kHz
DMSO:	7.1 kHz		Toluene:		5.4 kHz
Methanol:	7.0 kHz		Deuterium oxide:	6.3 kHz

- Shut off sample spinning if you want to run a NOESY experiment.

- Start the experiment by clicking the ZG button.
Check the duration of the experiment on the status line at the bottom of the window.
If it takes too long, press the red Stop button and decrease Points1D, Points2D and Acq.Points. (Try half the previous settings for each of the three parameters.)
If you have a dilute sample, increase Scans1D (multiples of 8).	
RESTART  ZG.
Don’t forget to save the experiment again after it has completed.	(Ctrl – S)

- Continue with 2D Processing (page 14).

## C.II	HETCOR (= HSQC, Duration with default settings approx. 1.5 hours)
 	HETCOR – long range (= HMBC, default settings approx. 10 hours)

- Press Ctrl – O and open a Proton parameter file.

- Record a proton spectrum, and process: Ctrl – B, Ctrl – F, Ctrl – H, calibrate and save it.

- Display the complete spectrum on the screen. 
Are there any signals above 9 ppm or below -1 ppm? (This is important for step 8!)

- Press Ctrl – O and open a Carbon-13 parameter file.

- Record a C-13 spectrum, Ctrl – B, Ctrl – E, Ctrl – F, Ctrl – H, calibrate and save it.

- Press Ctrl – O and open the file “HETCOR.tnt”, if you just want to see direct connectivities.
Choose “HETCOR-long range.tnt” instead, if you just are interested in long range couplings.
Use “File | Save as” to define the filename in your data directory.

- Select:	View | Dashboard and select the Acquisition tab.
Scans1D should be 32 for a normal HETCOR and 196 for a long range HETCOR.
By default, Points1D and Acq.Points are both 2048, and Points2D is 128.
These parameters determine the resolution and the duration of the experiments.

- If there were any CH proton signals above 9 ppm or below -1 ppm in your proton spectrum 
you will have to increase SW2D by 0.7 kHz for every additional ppm to avoid folding effects. 
OH and NH proton signals may be ignored!

- Switch to the Frequency tab. 
Make the F1 Base Freq. = 90.56 and the F2 Base Freq. = 360.13. Leave F1 Offset Freq. alone.
To avoid folding set the F2 Offset Freq. value according to your choice of solvent:
Acetone:	7.3 kHz		Methylene Chloride:	6.1 kHz
Acetonitrile:	7.6 kHz		Pyridine:		5.4 kHz
Benzene:	5.5 kHz		Trifluoro acetic acid:	4.7 kHz
Chloroform:	5.5 kHz		Tetrahydrofuran:	6.7 kHz
DMSO:	7.1 kHz		Toluene:		5.4 kHz
Methanol:	7.0 kHz		Deuterium oxide:	6.3 kHz

- Start the experiment by clicking the ZG button.
Check the duration of the experiment on the status line at the bottom of the window.
If it takes too long, press the red Stop button and decrease Points2D. 
(Try half the previous settings.)
If you have reserved more time, increase Scans1D (multiples of 32).

- Don’t forget to save the experiment again after it has completed.	(Ctrl – S)

- Continue with 2D Processing (page 14).

## Processing 2D spectra:

- First we have to process the reference spectra that you have recorded earlier:
Open the proton (and in the case of HETCOR also the carbon) reference spectrum and process them exactly the way you would usually do (proper phasing and calibration is required).

- Find a proton and a carbon signal that you know well and write down their chemical shifts.
These numbers are needed to calibrate the 2D spectra.

- Now you can open the 2-dimensional data file. Only the first trace will be displayed.
If a ppm axis appears instead of a time axis, advance to the last 2D trace and press Ctrl-F.

- Hold down the Shift key and click on the NDFT button to open the processing dialogue.

- Under the 1D tab check only: Use 1D settings and Fourier Transform.

- Under the 2D tab check only: Use 2D settings, Zero Fill, and Fourier Transform.

- Apodization:
For a COSY/NOESY in 1D and 2D choose Sin Bell with 30 for SB Shift and 1 for SB Skew.
  Set SB Width equal to the Points1D and Points2D values (listed above the spectrum).
For HETCOR use Exponential with LB=5 in dimension 1 and LB=3 in dimension 2.

- Click Do It when all parameters are correct.

- Select File | Save As and save the processed data with a different filename. 
(Suggestion: add a “P” to the filename.) Do not overwrite the original data file! 
Should anything go wrong during processing, you can recover by rerepeating from step 3.

If you prefer, you can open this file with the ACD-NMR program now.
It provides much better display and printing options than NTNMR.
Otherwise just proceed from here.

- Hold down the Shift key and click on the 2D display button.
 Select Magnitude and Contours, and select the appropriate aspect ratio: 
 Square for COSY/NOESY or None for HETCOR. Press Apply.
Select the Color: B/W and adjust the intensities by selecting suitable Start and End levels. 
Try to mazimize the signals and to minimize the noise.

- If a symmetrization is needed for your COSY or NOESY spectrum (not for HETCOR!!!):
 Select Commands | Symmetrize | Right to left, Low value.

- COSY/NOESY only: Right-click on the spectrum and select Reverse Data | Vertical. 
Adjust the Levels again to minimize the noise and to mazimize the signals.

- Double-click the horizontal axis at the bottom to open the 2D Axis Setup dialogue.
The Tick Marks settings should be: 1, 5, and 1. Check “Adjust Ticks on Zoom”. 
Switch to the vertical section by clicking the tab at the top and use the same settings.
HETCOR only: Change the vertical spectral width to the same value as in step 9.

- Right-click the diagonal peak or crosspeak that corresponds to the peaks selected in step 1.
Select Set Reference and click the “Enable Units to Frequency Domain” button, if present.
Then select ppm as units and enter the correct chemical shifts in both dimensions.

- Right-click in the spectrum and select Reset View. 
Select your region of interest and expand it using View | Zoom Toolbar.
First change the units to “ppm”, then enter the desired limits and click “Apply”. 
Did the signals disappear? Sometimes the proton limits have to be entered as negative numbers.

## Printing 2D spectra:

NTNMR’s Print Layout is extremely difficult to use with 2D spectra. Here is a better way to print:

- Stay in the 2D display and select Window | Cascade.

- Move the mouse pointer out of the spectrum to remove the crosshairs.

- Press the keyboard’s Print Screen key to copy the full screen to the clipboard.

- Start Microsoft Paint or use your favorite picture editing software.

- Paste the clipboard image into Paint (Ctrl-V).

- Activate the Select function and use the mouse to cut out the spectrum with its axes.

- Press Ctrl-C, then Ctrl-N, and then N. Enlarge the white area as much as possible.

- Press Ctrl-V to paste the 2D spectrum into Paint.

- Move the 2D spectrum to have 2 inches free space on the top and on the left side.

- Use File | Save As to save the image. (Monochrome Bitmap is sufficient.)

- Switch back to NTNMR and display the proton spectrum or the carbon spectrum.
Use Option | Amplitude Adjustment to expand exactly the same range as you did in the 2D.
Adjust the height to maximize the important signals without cutting them off.

- Shrink its window until the spectrum spans the same width as the 2D spectrum.
 Remember: In a HETCOR the proton axis is vertical. The C-13 axis is horizontal.

- Repeat from step 2.

- If you have a HETCOR you should have 3 bitmap files (the 2D, the 1H and the 13C)
A COSY and NOESY will only have 2 files (the 2D and the 1H projection)

- Finally, copy and paste the 2D spectrum and its 1D projections together into one image.
The 1D projections have to fit into the spaces above and at the left side (rotate) of the 2D.
Here is an example of how a HETCOR spectrum can be imported into a document:

# Emergency Shutdown

  - On the PC: Terminate any NMR Experiment:
(yellow or red STOP button in the NTNMR window)

  - Save the data to disk (Ctrl-S)

  - Press Ctrl – Alt – Del  and select Shutdown to shut down the Dell PC.

  - Switch off the Tecmag box (behind the console, left)

  - Pull the red “EMERGENCY” button on the front panel of the console.

Don’t forget to turn off the other spectrometers in the NMR lab. Now it it safe to leave.

# Restart

- Push the red EMERGENCY button back in.

- Go behind the console and switch the gray MAIN POWER RESET button to the left position.

- If the console starts beeping continuously, the SCM under the keyboard has to be reset.
follow the procedure on the printout on the console keyboard.
Set the suggested values for the SWEEP RATE and for the SWEEP AMPL.
Make sure all the buttons labeled “ON” are really on.

- Switch on the Tecmag box (behind the console, left)
(Six red and five green lights should come on.)

- Switch on the Dell PC.

- After the bootup has finished you should let the spectrometer warm up for at least 30 minutes before you start any experiments.

Important

When you are done with your session, please:

- Insert the water/ethanol sample

- Set the FIELD to 700 and the LOCK POWER to 35

- Set Z to 1220 and Z2 to 360

- Press the AUTO LOCK button

- Shut down the NTNMR Program and log off

- Sign the log book

- Remove all your items

- !!!!!!!!!!!!!!! CAP THE MAGNET !!!!!!!!!!!!!!!!!!

Thank You!

## Startup Parameters for the AM360 NMR Spectrometer (old values, do not use)

|   Only adjust these 3 settings:       |

|  | Settings on the Shim Panel | Proton NMR Parameters | Carbon-13 NMR Parameters |
| --- | --- | --- | --- |
| Solvent | FIELD | LOCK POWER | LOCK GAIN | F1 Freq. Offset | SR | F1 Freq. Offset | F2 Freq. Offset | SR |
| acetone-d6 | 4010 | 25* | 22 | 7478 | 5858.0 | 5217 | 7478 | -3840.5 |
| acetonitrile-d3 | 4050 | 30* | 22 | 7875 | 5900 | 5217 | 7475 | -3838.4 |
| benzene-d6 | 4360 | 25* | 22 | 5624 | 4003.6 | 4786 | 5624 | -4269.1 |
| chloroform-d | 4360 | 40* | 22 | 5750 | 3986.6 | 5217 | 5750 | -4239.0 |
| DMSO-d6 | 4010 | 25* | 22 | 7313 | 5692.7 | 5228 | 7313 | -3827.3 |
| methanol-d4 | 4100 | 30* | 22 | 7200 | 5692.0 | 5200 | 7200 |  |
| CD2Cl2 | 4250 | 30* | 22 | 5600 | 3800 | 5300 | 5600 |  |
| pyridine-d5 | 4500 | 30* | 22 | 5624 | 4003.6 | 4786 | 5624 | -4431.6 |
| TFA -d | 4800 | 45* | 22 | 4663 | 2504 | 4223 | 4663 | -4831.5 |
| THF-d8 | 4020 | 40* | 22 | 6927 | 5305 | 5228 | 6927 | -3989 |
| toluene-d8 | 4500 | 30* | 22 | 5625 | 4004.4 | 4786 | 5625 |  |
| water-d2 | 4170 | 30* | 22 | 6529 | 4908.5 | 4975 | 6529 | -4029.8 |

* For concentrated samples only (more than 10 % or more than 100 mg of sample.):  
 	Raise the LOCK POWER setting by 5 to 10 units

Sample Preparation:

1.    Choose an NMR solvent appropriate to your compound.

•  Chloroform is the standard solvent to try first. Be aware that it can be acidic.

•  If your compound is not soluble in chloroform, try benzene (nonpolar or average polarity compounds), acetone (dissolves almost anything), DMSO or methanol (polar compounds).

2.    Find a clean, dry NMR tube.

•  If your proton spectra consistently show a peak near 2.05 ppm, the tube wasn’t dry enough.
	(Traces of acetone may be present, even is the tube appears to be dry!)

•  After washing NMR tubes with acetone, lay the tubes flat on a layer of paper towels in the oven for at least two hours at max. 50 °C before using them again. A vacuum oven is preferred.

•  Standing the tubes upright in a flask or beaker in an oven is not recommended, as the tubes might warp or paramagnetic rust particles might fall into the tubes.

•  Compressed air contains mineral oils and is not recommended to dry NMR tubes.

•  Store clean NMR tubes uncapped, preferably laying down.

3.     Dissolve your sample.

•  Make sure your sample is free of solvent; if your compound is not volatile, 
placing the flask on a high vacuum line for 5-30 minutes is a good idea.

•  Measure the correct amount of sample into a vial.

•  For a solid, use 10-20 mg for a routine 1H and 50-100 mg for a routine 13C NMR. 
You can use less, but your experiments will take longer.

•  For an oil or liquid, dip a glass pipette into the liquid (one drop is enough for 1H NMR).

•  Dissolve the sample in 0.75 ml of the NMR solvent.

4.     De-gas your samples (for NOE or NOESY experiments).

•  Use several “Freeze – Pump – Thaw” cycles to remove dissolved oxygen.
	•  Easy alternative: bubble some N2 or Ar gas through your solution.

5.     Put the solution into the NMR tube. 
 	If any solid remains, filter the sample into the tube through a pipette with a cotton plug.

6.     Acquire your spectrum.